## Slide 1: Guidelines in Securing a License to Operate a COVID 19 Testing Laboratory in the Philippines
Doh ADMINISTRTIVE Order 2020-0014
Doh ADMINISTRTIVE Order 2020-0014-A
Doh ADMINISTRTIVE Order 2020-0014-B

## Slide 2
2 Types of Test done in a Licensed COVID 19 Testing Laboratory
A. Real Time Reverse Transcriptase Polymerase
Chain Reaction (rRT-PCR)
B. Cartridge Based Polymerase Chain Reaction

## Slide 3: Cartridge Based PCR vs rRT-PCR
Type of PCR Test | Specimen Source | Specimen Type | Turn Around Time | No. of Test Performed
Real Time Reverse Transcriptase – Polymerase Chain Reaction rRT-PCR | Nasopharyngeal swab Oropharyngeal swab | Inactivated Specimen | 24-48 Hours | Depends upon the machine (no. of wells, example 96 , 384)
Cartridge Based Polymerase Chain Reaction | Nasopharyngeal swab | Active Specimen | Approx. 45 Minutes | Depends upon the machine (no of modules, example 4, 16)

## Slide 4: Licensing Requirements for rRT PCR & Cartridge Based PCR
 | Personnel Requirement | Training Requirement
Real Time Reverse Transcriptase – Polymerase Chain Reaction rRT-PCR | Minimum number of personnel shall be 7, Shall be composed of the following: 1 Board Certified Clinical Pathologist 3 Full time Analyst per 8 hour shift shall be composed of EITHER 3 Registered Medical Technologist OR 2 Registered Medical Technologists and any allied health professionals with a Bachelor’s Degree relevant to the Job 1 Full-Time Laboratory Aide per 8 hour shift 1 Full Time Receptionist per 8 hour shift 1 Full Time Encoder per 8 hour shift | Training on Molecular Laboratory Diagnosis for Pathologist and RMTs Analyst Training on Biosafety and Biosecurity for RMTs Analyst, Lab Aide, Encoder and Receptionist
Rapid Cartridge Based Polymerase Chain Reaction Rapid PCR | 1 Board Certified Clinical Pathologist 1 RMT Analyst per 2 machines 1 Laboratory Aide, 1 Encoder per 4 Machines | Training on Cartridge Based Technology (Ex. GeneXpert System) for Pathologist and RMTs Analyst Training on Biosafety and Biosecurity for all personnel (Biosafety and Biosecurity training of Lab Aide and Encoder maybe in-house)

## Slide 5: Licensing Requirements for rRT PCR & Cartridge Based PCR
 | PHYSICAL FACILITY
Real Time Reverse Transcriptase – Polymerase Chain Reaction rRT-PCR | 1. Specimen Reception 2. Virus inactivation and nucleic acid extraction (pre PCR) 3. Reagent and Storage Handling 4. PCR and 5. Clerical Activities Prototype: DOH A.O. 2020-0014-A Annex D3 DOH A.O. 2020-0014-B Annex D1 Version3 See General Notes
Cartridge Based Polymerase Chain Reaction | 1. Specimen Reception 2. Specimen Processing and 3. Clerical Activities Prototype: DOH A.O. 2020-0014-A Annex D4, D6 DOH A.O. 2020-0014-B Annex D8 See General Notes

## Slide 6: Licensing Requirements for rRT PCR & Cartridge Based PCR
 | EQUIPMENT
rRT-PCR | Specimen Handling & Sample Specimen Receiving Room Template Adding Room Prep. Room 1. Autoclave 1. PCR Hood/Cabinet 1. Autoclave 2. BSC Class II A2 with dead air box 2. BSC Class II A2 3. Biomedical Ref. 2. Mini centrifuge 3. Biomedical Ref. w/ freezer (2) 3. Vortex Mixer 4. Cold rack for PCR Tubes Reagent Prep. Room 5. Heating Block 1. Biomedical Ref. w/ freezer + Biomedical Ref. 6. Micro centrifuge 2. Cold Rack for PCR Tube Amplification/PCR 7. Mini centrifuge 3. Micro centrifuge Room 8. Set of four volume 4. Mini centrifuge 1. Mini centrifuge/ micropipette w/ rack 5. Micropipette tips plate spinner 8. Micropipette tips 6. PCR Cabinet Laminar Flow 2. Real Time PCR 9. Vortex Mixer 7. Vortex Mixer machine 10. Computer w/ printer 8. Set of four volume micropipette with rack
Cartridge Based PCR | Biomedical Ref. for specimen Biomedical Ref for cartridge BSC Class II A2 Rapid PCR Machine with Kits and Cartridge

## Slide 8: STAGE 1
For Cartridge Based PCR COVID 19 Testing Lab
Letter of Intent – doh.covidlab.info@gmail.com
DOH HFSRB will send a self assessment tool designed by RITM
DOH Assessment Team will review the tool
For rRT-PCR COVID 19 Testing Lab
Letter of Intent – doh.covidlab.info@gmail.com
DOH will send a self assessment tool designed by RITM
DOH Assessment Team will review the tool
A. Apply for DOH – PTC
DOH HFSRB – Level 2 and Level 3 Hospitals
CHD RLED – Level 1 and Non Hospital Based
Transitory Provision : All PTC applications - HFSRB
B. DOH-License to Operate (DOH-LTO)
Filing of complete application requirements,
DOH HFSRB – Level 2 and Level 3 Hospitals
CHD RLED – Level 1 and Non Hospital Based
Transitory Provision : All LTO applications - HFSRB

## Slide 9
Complete documentary requirements shall consist of the following:
a. Notarized completely accomplished application form (ANNEX E DOH A.O. 2020-0014);
b. A copy of approved DOH-PTC and floor plan;
c. Notarized list of personnel, including photocopies of valid PRC identification card,
valid COVID- 19 proficiency training certificate from RITM, and copy of certificates of
all necessary trainings; (Annex A DOH A.O. 2020-0014)
d. List of equipment with specifications, reagents, and supplies; (Annex B DOH A.O. 2020-0014 )
e. Copy of Certificate of Product Registration (CPR) from Food and Drug Administration for
all equipment, reagents and supplies;
f. Accomplished Self-Assessment Tool for Licensing a COVID-19 testing laboratory;
Rapid PCR - Annex A2 DOH A.O. 2020-0014-B
rRT-PCR - Annex A1 DOH A.O. 2020-0014-B
g. For renewal, a copy of NEQAS, Certificate of Performance with PASSING results

## Slide 10: Stage 2
DOH A.O. 2020-0014
After evaluation of the submitted documents for technical completeness and correctness, the assigned team from HFSRB/CHD-RLED shall schedule an inspection date, in coordination with the team from RITM or its duly
recognized/authorized 3™ party assessors.
Applications shall be sent via e mail at doh.covidlab.info@gmail.com

## Slide 11: STAGE 3
PROOFS OF COMPLIANCE SHALL BE SENT TO:
doh.covidlab.info@gmail .com

## Slide 12: STAGE 4
Cartridge Based PCR
National Tuberculosis Reference Laboratory RITM
(Learning and Development Section)
- Didactic and Hands on Training on the use of
GeneXpert / Xpert Xpress Machine
- RITM
B. Real Time Reverse Transcriptase PCR (rRT-PCR)
- RITM will provide a PT panel that consists a set of
unknown samples. The laboratory technologists of the
applicant laboratory must then test these samples on PCR.
- The results of their testing will be the basis for determining
whether their facility and staff complement is capable to
correctly and accurately test for COVID-19.

## Slide 13: STAGE 5
RITM or its duly recognized/authorized 3" party assessors shall transmit to HFSRB/CHD-RLED a Letter of Recommendation when the facility is fully compliant to the standards and requirements of RITM.
The DOH-LTO a COVID-19 testing laboratory shall be issued only after full compliance to the standards and requirements by HFSRB/CHD-RLED and the RITM.
Validity of DOH – LTO
COVID 19 Testing Lab. using Cartridge Based Technology – 1 year
rRT-PCR COVID 19 Testing Laboratory – 1 year

## Slide 14
THANK YOU!!!